[image: image1.png]





CERTIFICATE OF CALIBRATION
	Cross Country Pipeline Supply Co. Inc. Certifies
	Serial No: 


	Model No: 780/785/PJM/DCPJM
	


has been calibrated and demonstrates the output voltage, as displayed on the dial of the detector, to be within +/- 5% of indicated setting over the range of 1000 -15000 Volts. Since output voltage is regulated, this tolerance is maintained regardless of load.

A Cross Country Pipeline Supply CO., Inc. Jeep Meter has been used to verify this performance.

In case of pulsed voltage, the Jeep Meter used has been calibrated with a high voltage generator (1-40 KV range) with a 5 micro second duration pulse administered as a rate of 27 HZ. This generator is periodically inspected every 6 months with a cathode ray oscilloscope and a precision high voltage probe. Both devices are in compliance with US Mil-Std-45662 and are routinely calibrated (approximately every six months) by an independent testing laboratory with traceability to the United States Nation Institute for Standards and Technology (NIST).
In the case of DC (direct current) a Jeep Meter is also used. Calibration of the Jeep Meter for DC is by way of a 1-5 KV DC power supply, which is in turn also routinely calibrated using the same oscilloscope and probe combination mentioned above.
Calibration certificates are on file and may be obtained by sending a written request to Cross Country Pipeline Supply, CO., Inc. It is suggested that this unit be checked for calibration at least once per year, and more frequently based on heavy usage.
To attain the performance certified here, a proper ground path and properly charged battery, as employed during this certification procedure, are essential and must be used.

It is suggested that this unit be checked for calibration at least once per year and more frequently based on usage.

Certified this 29TH DAY OF February 2023
By: (Signature)

       (Print) 
 Cross Country Pipeline Supply CO., Inc


